- PART 1 – GENERAL

  - References

    - All work shall be in conformance to the latest revision of « Quebec Building Code – Chapter I», unless otherwise indicated.

    - All work shall conform to the latest revised codes and standards that having jurisdiction, including but not limited to:

      - ANSI / ASME B40.2, Thermometers, Direct reading and Remote Reading;

      - CAN / CGSB-14.4, Thermometers, Liquid-in-Glass, Self-Indicating, Commercial / Industrial Type;

      - CAN / CGSB-14.5, Thermometers, Bimetallic, Self-Indicating, Commercial/Industrial Type.

  - Submittals

    - Product Data:

      - Submit shop drawings and technical data in conformance with client’s instructions.

    - Identify the features listed on the manufacturer's documentation:

      - Thermometer;

      - Thermowell.

  - Instruction and Maintenance Manual

    - Submit manufacturer’s installation and start-up instructions.

    - The maintenance manual will comprise of or indicate the following:

      - A description of the major components; the manufacturer, series or model reference number;

      - All details relating to the operation, care and maintenance of component;

      - A list of equivalent component replacements.

- PART 2 – PRODUCT

- Metal-Case, Industrial-Style, Liquid-in-Glass Thermometers:

- Basis-of-Design Product: Subject to compliance with requirements, provide Flo Fab Inc; Model FFBR9A or FFBRHW or comparable product by the following:

- Standard: ASME B40.200.

- Case: [Cast aluminum] <Insert material>; [7-inch (178-mm)] [9-inch (229-mm)] nominal size unless otherwise indicated.

- Case Form: [Adjustable angle] [Back angle] [Straight] <Insert form> unless otherwise indicated.

- Tube: Glass with magnifying lens and blue[ or red] organic liquid.

- Tube Background: Nonreflective aluminum with permanently etched scale markings graduated in [deg F (deg C)] [deg F and deg C].

- Window: [Glass] [or] [plastic] <Insert material>

- .

- Stem: [Aluminum] <Insert material> and of length to suit installation.

- Design for Thermowell Installation: Bare stem.

- Connector: 1-1/4 inches (32 mm), with ASME B1.1 screw threads.

- Accuracy: Plus or minus 1 percent of scale range or one scale division, to a maximum of 1.5 percent of scale range.

- Scale Range for Chilled-Water Piping: [0 to 120 deg F (Minus 20 to plus 49 deg C)].

- Scale Range for Heating, Hot-Water Piping: [32 to 240 deg F (0 to 150 deg C)].

- Scale Range for Low Temperature Hot-Water Piping: [32 to 180 deg F (0 to 115 deg C)].

- Scale Range for Domestic Water Piping: [32 to 180 deg F (0 to 115 deg C)].

  - Thermowell

    - Install copper or bronze thermowells in copper piping.

    - Install stainless steel or brass thermowells in steel piping.

- PART 3 – EXECUTION

  - Installation

    - Install thermometers at indicated locations and at the entrance and exit of these components:

      - Heat exchangers;

      - Heating and cooling coils;

      - Boilers;

      - Refrigeration;

      - Cooling towers;

      - Domestic hot water.

    - Use extension piping when the thermometers are mounted on insulated piping.

    - A thermowell should be used on all pressurized applications, to protect the stem of the thermometer from corrosion and physical damage, and to facilitate removal of the thermometer without disturbing the process.

    - Supply and pre-install fittings for thermometers where proper system balancing will require temperature measurements.

End of Section